## Input-Output Sheet
Enter room data (fill in blue boxes)
Building
Room
Fume hood present in room?
Room Height (ft)
Room Width (ft)
Room Length (ft)
Exhaust or supply in cfm - add fumehood exhaust to this
Total volume of cryogens stored (L) (including cryomagnets, traps, LN2 freezers, etc.)
Largest storage Dewar (L) (including cryomagnets, traps, LN2 freezers, etc.)
Largest transfer Dewar (L)
Routine work - please see comments for explanation of "Routine Work"
Number of 13 cu ft freezers with LN2 backup
Number of 25 cu ft freezers with LN2 backup
LN2 in cryomagnet (L) (enter the volume of N2 in Dewar)
LHe in cryomagnet (L) (enter the volume of He in Dewar)
Energy of cryomagnet (kJ) (enter 0 if unknown)
Calculated Values for Room - Do not enter values here!
Volume of Room (ft^3) - 15% for equipment | 0
Volume of Room (L) | 0
Quench Volume | 0
Exhaust or supply air in cfm | 0
Air exchanges per hour | #DIV/0!
Hazard Calculations
Routine operations - Do not enter values here!
| % O2 | Hazard Determination
Storage of cryogens | #DIV/0! | #DIV/0!
Routine work | #DIV/0! | #DIV/0!
Filling transfer Dewar | #DIV/0! | #DIV/0!
Refilling cryomagnet | #DIV/0! | #DIV/0!
Cool-down of cryomagnet | #DIV/0! | #DIV/0!
Nonroutine conditions (spills, accidents, power loss, etc.) - Do not enter values here
| % O2 | Hazard Determination
8 hr power outage (no air exchange) | #DIV/0! | #DIV/0!
Transfer Dewar spill (no air exchange) | #DIV/0! | #DIV/0!
Transfer Dewar spill (with air exchange after 1 hour) | #DIV/0! | #DIV/0!
Complete failure of storage dewar (no air exchange) | #DIV/0! | #DIV/0!
Complete failure of storage dewar (with air exchange - after 1 hr) | #DIV/0! | #DIV/0!
Magnet quench (no air exchange) | #DIV/0! | #DIV/0!
Magnet Quench (after 1 hour with air exchange) | #DIV/0! | #DIV/0!
Note on A18: Windows User: Routine work means - Work with cryogen that will actively be generating gas includes such tasks as working in open dewars (crystal mounting) or using liquid nitrogen to actively cool experiments (cryostream). All of the cryogen in the dewar is assumed to be used within one hour. Cryogen is assumed to be liquid helium for the sake of expansion ratio.

## Calculations (hidden)
Calculations
Enter room data (fill in blue boxes) |  |  | Symbol | Definition | Value | Units | Reference
|  |  | L | Liquid cryogen released / 24 hours | 0 | L
Building |  | 0 | Vr | Actual Room Volume (Room volume empty - 15% being full of equipment) | 0 | ft^3
Room |  | 0 | Vr | Actual Room Volume (Room volume empty - 15% being full of equipment) | 0 | L
Fume hood present in room? |  | 0 | E | expansion co-efficient | 757
|  |  | Ach | Air changes per hour | #DIV/0! | L
Room Height (ft) |  | 0 | Q | Volumetric flow | 0 | cfm
Room Width (ft) |  | 0 | Q | Hourly volumetric flow in liters/hour | 0 | L/hour
Room Length (ft) |  | 0 | Vlvt | Volume of Gaseous Cryogen released by venting total/ 24 hrs (in 24 hours) | 0 | L
|  |  | Vlvth | Volume of Gaseous Cryogen released by venting total/ hr (in 24 hours) | 0 | L
Exhaust or supply in cfm - add fumehood exhaust to this |  | 0 | Vlsd | Volume of largest storage dewar as an expanded gas | 0 | L
|  |  | Vltd | Volume of largest transfer dewar as an expanded gas | 0 | L
Total volume of cryogens stored (L) (including cryomagnets, traps, LN2 freezers, etc.) |  | 0 | Lr(13) | 13 cu ft freezer leak rate (l/hr) | 4 | L/hr
Largest storage Dewar (L) (including cryomagnets, traps, LN2 freezers, etc.) |  | 0 | Lr(25) | 25 cu ft freezer leak rate (l/hr) | 6.5 | L/hr
|  |  |  | Leak rate (% per day) | 5 | %/day
Largest transfer Dewar (L) |  | 0 |  | Cryo transfer loss (% per day) | 10 | %/day
|  |  |  | He Cool down loss (% of He Dewar) | 400
Routine work - please see comments for explanation of "Routine Work" |  | 0 |  | Nitrogen Cool down loss (% of N2 Dewar) | 300
Number of 13 cu ft freezers with LN2 backup |  | 0 |  | Normal Storage after 1 day | #DIV/0! | % cryogen
Number of 25 cu ft freezers with LN2 backup |  | 0 |  | Routine work - known volume of cryogen is consumed | #DIV/0! | % cryogen
|  |  |  | Filling a transfer dewar - assume 10% of total volume evaporates in 1 hr | #DIV/0! | % cryogen
LN2 in cryomagnet (L) (enter volume of N2 in Dewar) |  | 0 |  | Re-filling a cryogmagnet - assume 10% of total volume evaporates in 1 hr | #DIV/0! | % cryogen
LHe in cryomagnet (L) (enter volume of He in Dewar) |  | 0 |  | Cool-down of cryogmagnet - assume 20% for N2 loss, and 50% for He loss | #DIV/0! | % cryogen
Energy of cryomagnet (kJ) (enter 0 if unknown) |  | 0 |  | 8-hour power outage, no air exchange | #DIV/0! | % cryogen
|  |  |  | Transfer Dewar Spillage - entire contents with no air exchange | #DIV/0! | % cryogen
Calculated Values for Room |  |  |  | Transfer Dewar Spillage - entire contents with air exchange after 1 hr | #DIV/0! | % cryogen
|  |  |  | Complete failure of storage dewar (no air exchange) | #DIV/0! | % cryogen
Volume of Room (ft^3) - 15% for equip. | 0 |  |  | Complete failure of storage dewar (with air exchange - after 1 hr) | #DIV/0! | % cryogen
Volume of Room (L) | 0 |  |  | Magnetic Quench (no air exchange) | #DIV/0! | % cryogen
Quench Volume | 0 |  |  | Magnetic Quench (after 1 hour) | #DIV/0! | % cryogen
Exhaust or supply air in cfm | 0
Air exchanges per hour | #DIV/0!
Hazard Calculations
Routine operations
| % O2 | Hazard Determination
Storage of cryogens | #DIV/0! | #DIV/0!
Routine work | #DIV/0! | #DIV/0!
Filling transfer Dewar | #DIV/0! | #DIV/0!
Refilling cryomagnet | #DIV/0! | #DIV/0!
Cool-down of cryomagnet | #DIV/0! | #DIV/0!
Nonroutine conditions (spills, accidents, power loss, etc.)
| % O2 | Hazard Determination
8 hr power outage (no air exchange) | #DIV/0! | #DIV/0!
Transfer Dewar spill (no air exchange) | #DIV/0! | #DIV/0!
Transfer Dewar spill (with air exchange after 1 hour) | #DIV/0! | #DIV/0!
Complete failure of storage dewar (no air exchange) | #DIV/0! | #DIV/0!
Complete failure of storage dewar (with air exchange - after 1 hr) | #DIV/0! | #DIV/0!
Magnet quench (no air exchange) | #DIV/0! | #DIV/0!
Magnet Quench (after 1 hour with air exchange) | #DIV/0! | #DIV/0!
Note on A19: Windows User: Routine work means - Work with cryogen that will actively be generating gas includes such tasks as working in open dewars (crystal mounting) or using liquid nitrogen to actively cool experiments (cryostream). All of the cryogen in the dewar is assumed to be used within one hour. Cryogen is assumed to be liquid helium for the sake of expansion ratio.